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Technical Datasheet

VWR® 21X9CM HYGIENE FACEMASK WITH TIES
113-9162/9163

CHARACTERISTICS:

•	Made of non-woven fabric 20g TB/25g MB/28g PP
•	White colour
•	Ties
•	High bacterial filtration efficiency
•	Sterile (SAL 10-6) and non-sterile version
•	Natural rubber latex free
•	Meet IEST Level 2 (2,440 counts/ft^3 for ≥0.5um) for non-sterile version
•	Clean room package

MATERIAL SPECIFICATION

1. 20 g/m2 TB (Thermally Bonded Polypropylene) for Inner Layer
	Item
	Test Method
	Standard

	Weight (g/m2)
	GB/T4669-1995
	20+/-2

	Break Tensile Properties (N/5cm)
	MD
	GB/T3923.1-1997
	>50

	
	CD
	GB/T3923.1-1997
	>15

	Elongation Rate(%)
	MD
	GB/T3923.1-1997
	>30

	
	CD
	GB/T3923.1-1997
	>30


[bookmark: OLE_LINK1][bookmark: OLE_LINK2]2. 25 g/m2 MB (Melt Blown Polypropylene) for Middle Layer
	Item
	Test Method
	Standard

	
	
	Non-Sterile
	Sterile

	Weight(g/m2)
	GB/T4669-1995
	25+/-2

	Thickness(mm)
	ASTM D1777
	0.25+/-0.02

	Break Tensile Properties (N/5cm)
	MD
	GB/T3923.1-1997
	>10
	>10

	
	CD
	GB/T3923.1-1997
	>6.5
	>6.5

	Elongation Rate(%)
	MD
	GB/T3923.1-1997
	>10
	>8

	
	CD
	GB/T3923.1-1997
	>10
	>8

	Bacterial Filtration Efficiency (%)
	ASTM F2101-2014
	>99%
	>80%

	Delta P
	MIL-36945C
	<2.5
	<4


3. 28 g/m2 PP (Spunbonded Polypropylene) for Outer Layer
	Item
	Test Method
	Standard

	
	
	Non-Sterile
	Sterile

	Weight(g/m2)
	GB/T4669-1995
	28+/-2

	Thickness(mm)
	ASTM D1777
	0.28+/-0.02

	Break Tensile Properties (N/5cm)
	MD
	GB/T3923.1-1997
	>30
	>30

	
	CD
	GB/T3923.1-1997
	>23
	>23

	Elongation Rate(%)
	MD
	GB/T3923.1-1997
	>30
	>20

	
	CD
	GB/T3923.1-1997
	>30
	>20



PRODUCTS AND DIMENSION SPECIFICATION
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	ECN
	Sterile/Non Sterile
	BFE(%)
	Delta P
	Dimension

	
	
	
	
	A
	B
	C
	D
	E

	113-9162　
	Non Sterile
	>99
	<3.5
	210
	90
	120
	400
	10

	113-9163
	Sterile
	>80
	<3.5
	210
	90
	120
	400
	10

	Tolerance
	±5
	±5
	±5
	±10
	±5




PACKING:

	Package/Quantity
	1pcs/bag, 50bags/big bags, 10big bags/case

	Best Storage Conditions
	Temp : 0- 50°C, Humidity : 45-85% 
without sunshine exposure

	Country of Origin
	China



QUALITY ASSURANCE CRITERIA:

1. Dimension and Function - Critical (AQL1.5, S-2)
•	Product dimension out of spec.
•	Fabric characteristics out of spec

2. Visual Inspection/ Attributes - Major (AQL 2.5, L-II) 
•	Foreign Material 
•	Dirt 
•	Thread ends loose or too long(over 1cm) 
•	Holes/Tears 
•	Foul odor 
•	Discolor 
•	Cloth surface applied wrongly 

3. Visual Inspection/Attributes - Minor (AQL 4.0, L-II) 
•	Unwanted folds 
•	Elastic tightness non-consistent 



Disclaimer: VWR wants to state that this declaration will not discharge the user from his obligation to ensure the product is suitable for the intended use. -  If there is any doubt ask us.

Date of issue: Aug 2021
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